Table 1: Performance metrics for ARMIE sensor PM₂.₅ concentrations

Against reference instruments (SMPS + APS, DustTrak), stratified by source type and calibration method

| Reference | Method | Pearson´s r | MAE (µg/m³) | RMSE (µg/m³) | Mean Relative Error (%) | Mean Error (µg/m³) | LoA Low (µg/m³) | LoA High (µg/m³) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Candles |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.98 | 17.9 | 24.8 | -20.3897392 | -17.3 | -52.2 | 17.6 |
| DustTrak | Uncalibrated | 0.98 | 8.5 | 13.0 | -0.1822668 | -3.0 | -27.9 | 22.0 |
| SMPS + APS | Calibrated | 0.76 | 130.8 | 192.9 | -70.9733227 | -130.8 | -408.8 | 147.2 |
| SMPS + APS | Uncalibrated | 0.78 | 96.2 | 155.5 | -47.1360243 | -96.2 | -335.9 | 143.4 |
| Candles High RH (70 - 80%) |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.93 | 8.4 | 16.5 | 4.0942553 | -3.3 | -35.0 | 28.4 |
| DustTrak | Uncalibrated | 0.93 | 12.4 | 21.6 | 26.6081560 | 9.5 | -28.6 | 47.6 |
| SMPS + APS | Calibrated | 0.81 | 56.2 | 91.8 | -37.6677330 | -55.0 | -199.2 | 89.2 |
| SMPS + APS | Uncalibrated | 0.78 | 34.9 | 67.9 | 16.1629933 | -24.1 | -148.7 | 100.5 |
| Cigarette Smoke |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.96 | 10.1 | 16.1 | 32.0943423 | 8.9 | -17.6 | 35.3 |
| DustTrak | Uncalibrated | 0.96 | 13.9 | 22.7 | 33.4587678 | 13.4 | -22.5 | 49.3 |
| SMPS + APS | Calibrated | 0.78 | 69.4 | 83.2 | 396.6942937 | 69.4 | -20.9 | 159.7 |
| SMPS + APS | Uncalibrated | 0.78 | 47.2 | 62.5 | 211.4953349 | 44.3 | -42.3 | 130.9 |
| Cooking |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.95 | 20.0 | 29.6 | 19.4611508 | 17.9 | -28.3 | 64.1 |
| DustTrak | Uncalibrated | 0.96 | 25.4 | 35.4 | 25.4767667 | 24.3 | -26.4 | 74.9 |
| SMPS + APS | Calibrated | 0.90 | 162.3 | 192.6 | 267.7259201 | 160.9 | -47.1 | 368.9 |
| SMPS + APS | Uncalibrated | 0.92 | 65.0 | 77.6 | 104.4740477 | 62.3 | -28.3 | 152.9 |
| Cooking High RH (70 - 80%) |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.97 | 6.4 | 10.1 | 15.5405746 | 1.6 | -18.1 | 21.2 |
| DustTrak | Uncalibrated | 0.97 | 11.8 | 18.6 | 25.4985599 | 9.6 | -21.6 | 40.8 |
| SMPS + APS | Calibrated | 0.91 | 45.7 | 59.6 | 164.7314142 | 43.5 | -36.3 | 123.4 |
| SMPS + APS | Uncalibrated | 0.92 | 30.9 | 39.6 | 117.1147672 | 28.3 | -26.1 | 82.8 |