Table 1: Table 3: Performance metrics based on the leave-one-source-out approach

| Comparison | Method | Pearson's r | Mean Relative Error (%) | Mean Error (µg/m³) | LoA Low (µg/m³) | LoA High (µg/m³) |
| --- | --- | --- | --- | --- | --- | --- |
| Candles |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.98 | -20.45 | -17.26 | -51.99 | 17.47 |
| DustTrak | Uncalibrated | 0.98 | -0.18 | -2.96 | -27.89 | 21.98 |
| SMPS + APS | Calibrated | 0.78 | -76.76 | -124.08 | -384.37 | 136.22 |
| SMPS + APS | Uncalibrated | 0.78 | -47.14 | -96.24 | -335.85 | 143.38 |
| Candles High RH (70 - 80%) |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.93 | 3.69 | -1.3 | -33.78 | 31.18 |
| DustTrak | Uncalibrated | 0.93 | 26.61 | 9.49 | -28.63 | 47.61 |
| SMPS + APS | Calibrated | 0.8 | -35.81 | -40.38 | -166.38 | 85.62 |
| SMPS + APS | Uncalibrated | 0.78 | 16.16 | -24.06 | -148.66 | 100.54 |
| Cigarette Smoke |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.96 | 37.07 | 8.88 | -16.02 | 33.79 |
| DustTrak | Uncalibrated | 0.96 | 33.46 | 13.42 | -22.47 | 49.32 |
| SMPS + APS | Calibrated | 0.78 | 398.3 | 69.28 | -19.65 | 158.21 |
| SMPS + APS | Uncalibrated | 0.78 | 211.5 | 44.27 | -42.34 | 130.88 |
| Cooking |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.95 | 13.64 | 10.93 | -29.34 | 51.2 |
| DustTrak | Uncalibrated | 0.96 | 25.48 | 24.26 | -26.38 | 74.9 |
| SMPS + APS | Calibrated | 0.92 | 223.53 | 129.31 | -26.43 | 285.05 |
| SMPS + APS | Uncalibrated | 0.92 | 104.47 | 62.33 | -28.27 | 152.92 |
| Cooking High RH (70 - 80%) |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.97 | 2.38 | 1.17 | -20.02 | 22.37 |
| DustTrak | Uncalibrated | 0.97 | 25.5 | 9.57 | -21.6 | 40.75 |
| SMPS + APS | Calibrated | 0.92 | 194.86 | 42.16 | -16.95 | 101.27 |
| SMPS + APS | Uncalibrated | 0.92 | 117.11 | 28.31 | -26.12 | 82.75 |