Table 1: Performance metrics for ARMIE sensor PM₂.₅ concentrations

Against reference instruments (SMPS + APS, DustTrak), stratified by source type and calibration method

| Reference | Method | Pearson´s r | MAE (µg/m³) | RMSE (µg/m³) | Mean Relative Error (%) | Mean Error (µg/m³) | LoA Low (µg/m³) | LoA High (µg/m³) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Candles |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.93 | 5.6 | 7.5 | -18.297427 | -5.4 | -15.6 | 4.9 |
| DustTrak | Uncalibrated | 0.93 | 3.5 | 4.9 | 2.494201 | -0.4 | -10.0 | 9.2 |
| SMPS + APS | Calibrated | 0.53 | 47.8 | 91.0 | -70.818340 | -47.8 | -200.0 | 104.3 |
| SMPS + APS | Uncalibrated | 0.52 | 35.2 | 83.3 | -39.324805 | -35.2 | -183.5 | 113.1 |
| Candles High RH (70 − 80%) |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.79 | 3.2 | 9.0 | 6.134705 | -0.3 | -17.9 | 17.3 |
| DustTrak | Uncalibrated | 0.78 | 4.7 | 9.5 | 24.175428 | 2.6 | -15.2 | 20.5 |
| SMPS + APS | Calibrated | 0.57 | 9.9 | 24.0 | -27.318663 | -9.0 | -52.8 | 34.7 |
| SMPS + APS | Uncalibrated | 0.49 | 9.1 | 23.1 | 37.421238 | 0.7 | -44.7 | 46.0 |
| Cigarette Smoke |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.93 | 3.9 | 5.3 | 26.888928 | 3.2 | -5.1 | 11.6 |
| DustTrak | Uncalibrated | 0.93 | 4.5 | 5.8 | 30.758901 | 4.0 | -4.4 | 12.4 |
| SMPS + APS | Calibrated | 0.64 | 50.0 | 53.2 | 506.873578 | 50.0 | 14.2 | 85.8 |
| SMPS + APS | Uncalibrated | 0.69 | 21.2 | 22.2 | 222.833000 | 21.2 | 8.4 | 34.1 |
| Cooking |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.32 | 9.3 | 12.6 | 18.402751 | 4.1 | -19.6 | 27.8 |
| DustTrak | Uncalibrated | 0.33 | 10.3 | 13.0 | 22.428953 | 5.4 | -18.0 | 28.9 |
| SMPS + APS | Calibrated | -0.10 | 51.3 | 55.7 | 245.116713 | 49.1 | -2.8 | 101.1 |
| SMPS + APS | Uncalibrated | -0.04 | 19.7 | 22.9 | 78.866039 | 12.6 | -25.3 | 50.6 |
| Cooking High RH (70 − 80%) |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.82 | 4.1 | 7.8 | 11.760444 | 0.4 | -14.9 | 15.7 |
| DustTrak | Uncalibrated | 0.83 | 5.8 | 8.4 | 27.226335 | 3.1 | -12.3 | 18.5 |
| SMPS + APS | Calibrated | 0.32 | 15.3 | 18.8 | 136.150839 | 12.0 | -16.5 | 40.5 |
| SMPS + APS | Uncalibrated | 0.41 | 17.1 | 19.6 | 158.214947 | 14.8 | -10.4 | 40.0 |