Table 1: Performance metrics for ARMIE sensor PM₂.₅ concentrations

Against reference instruments (SMPS + APS, DustTrak), stratified by sensor node and calibration method

| Reference | Method | r | MAE (µg/m³) | RMSE | Mean Relative Error (%) | Mean Error (µg/m³) | LoA Low (µg/m³) | LoA High (µg/m³) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.87 | 3.9 | 7.3 | 5.923406 | -0.2 | -14.7 | 14.2 |
| DustTrak | Uncalibrated | 0.87 | 5.4 | 7.7 | 21.110589 | 2.4 | -12.0 | 16.9 |
| SMPS + APS | Calibrated | 0.40 | 17.3 | 37.5 | 65.424267 | 0.6 | -73.4 | 74.7 |
| SMPS + APS | Uncalibrated | 0.29 | 18.8 | 38.6 | 73.558989 | -0.9 | -77.1 | 75.3 |
| 2 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.66 | 6.2 | 13.1 | 2.601445 | -2.7 | -28.0 | 22.6 |
| DustTrak | Uncalibrated | 0.65 | 7.0 | 13.0 | 19.591375 | 0.1 | -25.5 | 25.7 |
| SMPS + APS | Calibrated | 0.42 | 23.5 | 57.3 | 41.699420 | -9.2 | -120.6 | 102.2 |
| SMPS + APS | Uncalibrated | 0.30 | 26.3 | 60.8 | 49.616228 | -12.4 | -129.7 | 104.9 |
| 3 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.91 | 3.1 | 5.6 | 18.503397 | 2.2 | -7.9 | 12.3 |
| DustTrak | Uncalibrated | 0.90 | 4.8 | 7.0 | 33.462780 | 4.3 | -6.6 | 15.3 |
| SMPS + APS | Calibrated | 0.40 | 19.8 | 52.7 | 78.848365 | 2.9 | -100.8 | 106.7 |
| SMPS + APS | Uncalibrated | 0.30 | 20.8 | 55.1 | 70.627963 | -3.5 | -112.0 | 105.0 |
| 4 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.80 | 4.7 | 10.7 | -6.604512 | -3.2 | -23.4 | 17.0 |
| DustTrak | Uncalibrated | 0.81 | 5.0 | 9.9 | 9.026908 | 0.0 | -19.6 | 19.5 |
| SMPS + APS | Calibrated | 0.42 | 23.3 | 58.3 | 37.340075 | -9.2 | -122.7 | 104.2 |
| SMPS + APS | Uncalibrated | 0.34 | 23.1 | 60.6 | 52.210854 | -8.3 | -126.8 | 110.1 |
| 5 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.96 | 2.4 | 3.7 | 6.279967 | 0.9 | -6.1 | 7.9 |
| DustTrak | Uncalibrated | 0.95 | 3.9 | 5.1 | 20.718930 | 3.4 | -4.0 | 10.9 |
| SMPS + APS | Calibrated | 0.47 | 18.0 | 52.6 | 36.106698 | -5.1 | -108.3 | 98.1 |
| SMPS + APS | Uncalibrated | 0.37 | 20.8 | 55.8 | 52.752019 | -5.9 | -115.2 | 103.4 |
| 6 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.97 | 2.5 | 3.6 | -5.841204 | -1.4 | -8.0 | 5.1 |
| DustTrak | Uncalibrated | 0.97 | 2.5 | 3.6 | 8.884581 | 1.2 | -5.5 | 7.9 |
| SMPS + APS | Calibrated | 0.42 | 17.9 | 49.2 | 36.502656 | -2.5 | -99.3 | 94.4 |
| SMPS + APS | Uncalibrated | 0.35 | 18.8 | 51.3 | 45.233892 | -4.6 | -105.3 | 96.0 |
| 7 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.87 | 4.1 | 6.4 | 13.028989 | 1.2 | -11.3 | 13.6 |
| DustTrak | Uncalibrated | 0.86 | 5.4 | 7.4 | 27.615347 | 3.4 | -9.5 | 16.3 |
| SMPS + APS | Calibrated | 0.53 | 15.5 | 24.0 | 80.760394 | 6.9 | -38.4 | 52.3 |
| SMPS + APS | Uncalibrated | 0.36 | 17.1 | 21.8 | 64.247114 | -0.7 | -43.7 | 42.3 |
| 8 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.87 | 4.5 | 7.5 | 10.858994 | 0.8 | -13.9 | 15.5 |
| DustTrak | Uncalibrated | 0.87 | 5.8 | 8.2 | 24.883548 | 3.2 | -11.6 | 18.0 |
| SMPS + APS | Calibrated | 0.52 | 16.5 | 29.5 | 76.953333 | 5.3 | -51.8 | 62.5 |
| SMPS + APS | Uncalibrated | 0.44 | 18.6 | 28.7 | 63.515188 | -1.8 | -58.3 | 54.7 |
| 9 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.94 | 3.1 | 4.5 | 10.470744 | 1.4 | -7.0 | 9.7 |
| DustTrak | Uncalibrated | 0.94 | 4.6 | 5.7 | 24.233671 | 3.7 | -4.8 | 12.2 |
| SMPS + APS | Calibrated | 0.48 | 18.0 | 50.0 | 66.872950 | 1.8 | -96.7 | 100.3 |
| SMPS + APS | Uncalibrated | 0.34 | 20.2 | 54.0 | 55.443677 | -4.7 | -110.7 | 101.3 |