Table 1: Table 1: Performance metrics based on the leave-one-sensor-out approach

| Reference | Method | Pearson's r | Mean Relative Error (%) | Mean Error (µg/m³) | LoA Low (µg/m³) | LoA High (µg/m³) |
| --- | --- | --- | --- | --- | --- | --- |
| DustTrak | Calibrated | 0.84 | 5.95 | -0.16 | -15.12 | 14.79 |
| DustTrak | Uncalibrated | 0.85 | 20.88 | 2.38 | -12.58 | 17.35 |
| SMPS + APS | Calibrated | 0.41 | 56.65 | -1.21 | -94.75 | 92.33 |
| SMPS + APS | Uncalibrated | 0.33 | 57.81 | -5 | -102.44 | 92.44 |