Table 1: Performance metrics for ARMIE sensor PM₂.₅ concentrations

Against reference instruments (SMPS + APS, DustTrak), stratified by sensor node and calibration method

| Reference | Method | r | MAE (µg/m³) | RMSE | Mean Relative Error (%) | Mean Error (µg/m³) | LoA Low (µg/m³) | LoA High (µg/m³) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.98 | 7.4 | 11.1 | 5.471660 | 0.1 | -21.7 | 21.9 |
| DustTrak | Uncalibrated | 0.98 | 13.7 | 19.8 | 20.743746 | 11.0 | -21.3 | 43.3 |
| SMPS + APS | Calibrated | 0.53 | 45.9 | 74.0 | 83.467546 | 10.7 | -133.2 | 154.5 |
| SMPS + APS | Uncalibrated | 0.54 | 44.2 | 71.9 | 65.923933 | 5.3 | -135.7 | 146.3 |
| 2 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.94 | 11.2 | 20.2 | 3.375215 | -6.1 | -44.0 | 31.8 |
| DustTrak | Uncalibrated | 0.94 | 13.8 | 20.8 | 19.011399 | 4.0 | -36.1 | 44.1 |
| SMPS + APS | Calibrated | 0.50 | 60.1 | 101.8 | 65.694826 | -16.0 | -213.5 | 181.5 |
| SMPS + APS | Uncalibrated | 0.49 | 60.8 | 103.5 | 48.585014 | -21.5 | -220.5 | 177.4 |
| 3 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.95 | 9.9 | 19.2 | 20.917273 | 4.1 | -32.8 | 41.1 |
| DustTrak | Uncalibrated | 0.94 | 16.1 | 27.9 | 36.368455 | 14.4 | -32.6 | 61.3 |
| SMPS + APS | Calibrated | 0.53 | 61.0 | 100.8 | 94.803680 | -2.8 | -200.7 | 195.2 |
| SMPS + APS | Uncalibrated | 0.52 | 60.8 | 102.0 | 69.375313 | -10.8 | -210.1 | 188.5 |
| 4 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.95 | 9.3 | 19.2 | -5.221074 | -6.7 | -42.0 | 28.5 |
| DustTrak | Uncalibrated | 0.94 | 12.3 | 20.1 | 9.526248 | 4.2 | -34.4 | 42.8 |
| SMPS + APS | Calibrated | 0.54 | 53.2 | 90.2 | 58.782100 | -8.1 | -184.7 | 168.5 |
| SMPS + APS | Uncalibrated | 0.53 | 52.4 | 91.0 | 44.132294 | -12.6 | -189.7 | 164.5 |
| 5 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.95 | 9.2 | 16.2 | 3.722185 | -2.0 | -33.6 | 29.7 |
| DustTrak | Uncalibrated | 0.95 | 12.3 | 21.3 | 19.234504 | 9.4 | -28.3 | 47.0 |
| SMPS + APS | Calibrated | 0.52 | 56.6 | 98.9 | 59.220007 | -10.7 | -203.8 | 182.3 |
| SMPS + APS | Uncalibrated | 0.51 | 57.4 | 100.1 | 47.834584 | -14.8 | -209.2 | 179.6 |
| 6 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.95 | 9.9 | 17.4 | -1.380677 | -3.7 | -37.2 | 29.7 |
| DustTrak | Uncalibrated | 0.95 | 9.4 | 20.4 | 12.192225 | 6.1 | -32.1 | 44.4 |
| SMPS + APS | Calibrated | 0.52 | 54.7 | 98.1 | 57.917424 | -9.9 | -201.7 | 181.8 |
| SMPS + APS | Uncalibrated | 0.52 | 55.1 | 99.4 | 41.588580 | -15.4 | -208.3 | 177.6 |
| 7 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.96 | 10.6 | 14.8 | 9.097278 | 0.6 | -28.4 | 29.6 |
| DustTrak | Uncalibrated | 0.96 | 14.1 | 21.2 | 22.395071 | 10.6 | -25.6 | 46.7 |
| SMPS + APS | Calibrated | 0.52 | 63.1 | 101.8 | 79.709459 | -7.6 | -207.1 | 191.8 |
| SMPS + APS | Uncalibrated | 0.51 | 63.6 | 103.4 | 52.494561 | -16.1 | -216.8 | 184.5 |
| 8 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.97 | 10.5 | 15.1 | 10.610295 | 2.7 | -26.4 | 31.9 |
| DustTrak | Uncalibrated | 0.96 | 16.4 | 24.0 | 24.128319 | 12.8 | -27.0 | 52.6 |
| SMPS + APS | Calibrated | 0.53 | 59.1 | 98.9 | 80.833437 | -3.9 | -198.0 | 190.2 |
| SMPS + APS | Uncalibrated | 0.52 | 59.2 | 100.2 | 55.165104 | -12.2 | -207.6 | 183.2 |
| 9 |  |  |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.96 | 9.8 | 16.0 | 7.548248 | 0.7 | -30.7 | 32.1 |
| DustTrak | Uncalibrated | 0.95 | 13.1 | 23.1 | 21.128218 | 10.7 | -29.7 | 51.0 |
| SMPS + APS | Calibrated | 0.52 | 58.8 | 97.9 | 76.255190 | -5.0 | -196.9 | 187.0 |
| SMPS + APS | Uncalibrated | 0.51 | 59.1 | 99.4 | 49.060988 | -13.4 | -206.8 | 179.9 |