Table 1: Table 3: Performance metrics based on the leave-one-source-out approach

| Comparison | Method | Pearson's r | Mean Relative Error (%) | Mean Error (µg/m³) | LoA Low (µg/m³) | LoA High (µg/m³) |
| --- | --- | --- | --- | --- | --- | --- |
| Candles |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.93 | -18.65 | -5.26 | -15.38 | 4.86 |
| DustTrak | Uncalibrated | 0.93 | 2.49 | -0.43 | -10.02 | 9.16 |
| SMPS + APS | Calibrated | 0.53 | -72.72 | -47.18 | -197.93 | 103.57 |
| SMPS + APS | Uncalibrated | 0.52 | -39.32 | -35.21 | -183.5 | 113.09 |
| Candles High RH (70 - 80%) |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.79 | 5.7 | -0.18 | -17.84 | 17.48 |
| DustTrak | Uncalibrated | 0.78 | 24.18 | 2.63 | -15.22 | 20.49 |
| SMPS + APS | Calibrated | 0.53 | -20.15 | -4.16 | -49.08 | 40.77 |
| SMPS + APS | Uncalibrated | 0.49 | 37.42 | 0.65 | -44.72 | 46.02 |
| Cigarette Smoke |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.93 | 30.91 | 3.47 | -4.77 | 11.72 |
| DustTrak | Uncalibrated | 0.93 | 30.76 | 3.98 | -4.41 | 12.37 |
| SMPS + APS | Calibrated | 0.63 | 488.18 | 46.22 | 26.56 | 65.87 |
| SMPS + APS | Uncalibrated | 0.69 | 222.83 | 21.25 | 8.37 | 34.13 |
| Cooking |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.33 | 15.02 | 2.97 | -20.42 | 26.36 |
| DustTrak | Uncalibrated | 0.33 | 22.43 | 5.42 | -18.01 | 28.85 |
| SMPS + APS | Calibrated | -0.07 | 175.69 | 34.04 | -8.17 | 76.26 |
| SMPS + APS | Uncalibrated | -0.04 | 78.87 | 12.63 | -25.32 | 50.59 |
| Cooking High RH (70 - 80%) |  |  |  |  |  |  |
| DustTrak | Calibrated | 0.82 | 6.73 | 0.3 | -14.96 | 15.57 |
| DustTrak | Uncalibrated | 0.83 | 27.23 | 3.12 | -12.25 | 18.5 |
| SMPS + APS | Calibrated | 0.4 | 157.87 | 17.02 | -19.45 | 53.49 |
| SMPS + APS | Uncalibrated | 0.41 | 158.21 | 14.78 | -10.42 | 39.98 |